**Integrating the Healthcare Enterprise**



**IHE Radiology**

**Technical Framework Supplement**

**Standardized Operational Log of Events**

**(SOLE)**

**Draft in preparation for Public Comment**

Date: February 24, 2017

Author: Rob Horn

Email: <domain\_name@ihe.net>

**Please verify you have the most recent version of this document.** See [here](http://ihe.net/Technical_Frameworks/) for Trial Implementation and Final Text versions and [here](http://ihe.net/Public_Comment/) for Public Comment versions.

**Foreword**

This is a supplement to the IHE Radiology Technical Framework <VX.X>. Each supplement undergoes a process of public comment and trial implementation before being incorporated into the volumes of the Technical Frameworks.

General information about IHE can be found at: [www.ihe.net](http://www.ihe.net/).

Information about the IHE Radiology domain can be found at: [ihe.net/IHE\_Domains](D:\\Users\\mopoo\\AppData\\Roaming\\Microsoft\\Word\\ihe.net\\IHE_Domains\\).

Information about the organization of IHE Technical Frameworks and Supplements and the process used to create them can be found at: [http://ihe.net/IHE\_Process](http://ihe.net/IHE_Process/) and [http://ihe.net/Profiles](http://ihe.net/Profiles/).

The current version of the IHE RadiologyTechnical Framework can be found at: [http://ihe.net/Technical\_Frameworks](http://ihe.net/Technical_Frameworks/).

Comments may be submitted on IHE Technical Framework templates any time at [http://ihe.net/Templates\_Public\_Comments](http://ihe.net/Templates_Public_Comments/). Please enter comments/issues as soon as they are found. Do not wait until a future review cycle is announced.

CONTENTS

[Introduction to this Supplement 6](#_Toc476134887)

[Open Issues and Questions 7](#_Toc476134888)

[Closed Issues 7](#_Toc476134889)

[General Introduction 11](#_Toc476134890)

[Appendix A – Actor Summary Definitions 11](#_Toc476134891)

[Appendix B – Transaction Summary Definitions 11](#_Toc476134892)

[Glossary 11](#_Toc476134893)

[Volume 1 – Profiles 12](#_Toc476134894)

[<](#_Toc476134895)*[Copyright Licenses>](#_Toc476134895)* [12](#_Toc476134895)

[<](#_Toc476134896)*[Domain-specific additions>](#_Toc476134896)* [12](#_Toc476134896)

[X Standardized Operational Log of Events (SOLE) Profile 13](#_Toc476134897)

[X.1 SOLE Actors, Transactions, and Content Modules 13](#_Toc476134898)

[X.1.1 Actor Descriptions and Actor Profile Requirements 15](#_Toc476134899)

[X.1.1.1 Event Reporter 15](#_Toc476134900)

[X.1.1.2 Event Repository 15](#_Toc476134901)

[X.1.1.2 Event Consumer 15](#_Toc476134902)

[X.2 SOLE Actor Options 16](#_Toc476134903)

[X.2.1 Send Multiple Event Option 16](#_Toc476134904)

[X.2.2 Receive Multiple Events Option 16](#_Toc476134905)

[X.2.3 Retrieve ATNA Audit Event Option 16](#_Toc476134906)

[X.3 SOLE Required Actor Groupings 16](#_Toc476134907)

[X.4 SOLE Overview 17](#_Toc476134908)

[X.4.1 Concepts 17](#_Toc476134909)

[X.4.1.0 Events and extending events. 17](#_Toc476134910)

[X.4.1.2 Query 17](#_Toc476134911)

[X.4.1.2.1 Retrieve Syslog Message (ITI-82) 18](#_Toc476134912)

[X.4.1.2.2 Retrieve ATNA Audit Record (ITI-81) 18](#_Toc476134913)

[X.4.1.3 Filter and Forward 19](#_Toc476134914)

[X.4.1.4 RESTful Delivery (mobile) 19](#_Toc476134915)

[X.4.2 Use Cases 19](#_Toc476134916)

[X.4.2.1 Use Case #1: Track Study Reading Activities 20](#_Toc476134917)

[X.4.2.2 SOLE Event Analysis 22](#_Toc476134918)

[X.4.2.3 Delayed Event Delivery (mobile) 23](#_Toc476134919)

[X.4.2.4 Dashboard 24](#_Toc476134920)

[X.4.1.4 RESTful Delivery (outside analysis) 25](#_Toc476134921)

[X.4.3 Contents of SOLE messages 26](#_Toc476134922)

[X.5 SOLE Security Considerations 26](#_Toc476134923)

[X.5.1 Security Considerations for Actors 26](#_Toc476134924)

[X.5.2 Security Considerations for Event Reports 26](#_Toc476134925)

[X.6 SOLE Cross Profile Considerations 26](#_Toc476134926)

[Volume 2 – Transactions 28](#_Toc476134927)

[3.Y Transfer Multiple Event Reports [RAD-XX] 28](#_Toc476134928)

[3.Y.1 Scope 28](#_Toc476134929)

[3.Y.2 Actor Roles 28](#_Toc476134930)

[3.Y.3 Referenced Standards 29](#_Toc476134931)

[3.Y.4 Interaction Diagram 29](#_Toc476134932)

[3.Y.4.2.1 Trigger Events 29](#_Toc476134933)

[3.Y.4.2.2 Message Semantics 29](#_Toc476134934)

[3.Y.4.2.2.1 Resources 30](#_Toc476134935)

[3.Y.4.2.2.2 Query Parameters 30](#_Toc476134936)

[3.Y.4.2.2.3 Request Header Fields 30](#_Toc476134937)

[3.Y.4.2.2.4 Request Payload 30](#_Toc476134938)

[3.Y.4.2.2.4.1 JSON encoded array of Syslog Messages (Example) 30](#_Toc476134939)

[3.Y.4.2.3 Behavior 31](#_Toc476134940)

[3.Y.4.2.4 Response 31](#_Toc476134941)

[3.Y.4.2.4.1 Status Codes 32](#_Toc476134942)

[3.Y.4.2.4.2 Response Header Fields 32](#_Toc476134943)

[3.Y.4.2.4.3 Response Payload 32](#_Toc476134944)

[3.82.4.2.3 Expected Actions 32](#_Toc476134945)

[3.Y.5 Security Considerations 33](#_Toc476134946)

[3.Y.5.1 Security Audit Considerations 33](#_Toc476134947)

[3.Y.5.1.(z) <Actor> Specific Security Considerations 33](#_Toc476134948)

[Appendices 34](#_Toc476134949)

[Appendix A – <Appendix A Title> 34](#_Toc476134950)

[A.1 <Add Title> 34](#_Toc476134951)

[Appendix B – <Appendix B Title> 34](#_Toc476134952)

[B.1 <Add Title> 34](#_Toc476134953)

[Volume 2 Namespace Additions 34](#_Toc476134954)

[Volume 3 – Content Modules 35](#_Toc476134955)

[5 Namespaces and Vocabularies 36](#_Toc476134956)

[6 Content Modules 38](#_Toc476134957)

[6.X SOLE Event Definitions 38](#_Toc476134958)

[6.X.1 SWIM and SOLE Event selection 38](#_Toc476134959)

[6.X.2 Event Semantics 38](#_Toc476134960)

[6.X.3 Event Reports. 40](#_Toc476134961)

[6.x.3.1 OrderEntered 40](#_Toc476134962)

[6.x.3.2 AppointmentTimeScheduled 41](#_Toc476134963)

[6.x.3.3 PatientArrived 41](#_Toc476134964)

[6.x.3.4 RoomAssigned 41](#_Toc476134965)

[6.x.3.5 PatientIn 42](#_Toc476134966)

[6.x.3.6 PatientOut (RID45899) 42](#_Toc476134967)

[6.x.3.7 FirstImageCollected (RID46000) 42](#_Toc476134968)

[6.x.3.8 ImagingComplete (RID45835) 42](#_Toc476134969)

[6.x.3.9 StudyPrepared (RID45914) 43](#_Toc476134970)

[6.x.3.10 QCCompleted (RID28816) 43](#_Toc476134971)

[6.x.3.11 ReportDictated (RID45859) 43](#_Toc476134972)

[6.x.3.12 ReportApproved (RID45924) 44](#_Toc476134973)

[6.x.3.13 ReportPublished (RID45865) 44](#_Toc476134974)

[6.x.3.14 ExamArchiveCommit (RID4580) 44](#_Toc476134975)

[6.x.3.15 ExamTransferToBilling (RID45836) 45](#_Toc476134976)

[6.x.3.16 PatientMerged (RID45898) 45](#_Toc476134977)

[6.x.3.17 ExamReassigned (RID45863) 45](#_Toc476134978)

[6.x.3.18 ExamRemoved (RID45856) 45](#_Toc476134979)

[6.x.3.19 ExamOpenedFoReporting (RID45893) 46](#_Toc476134980)

[6.x.3.20 ExamCancelled (RID45862) 46](#_Toc476134981)

[6.x.3.21 ExamExceptionDetected (RAD-xx01) 46](#_Toc476134982)

[6.x.3.22 ExamExceptionResolved (RAD-xx02) 47](#_Toc476134983)

[6.x.3.23 ExamDeleted (RAD-xx03) 47](#_Toc476134984)

[6.x.3.23 CritNotification (RID45854) 47](#_Toc476134985)

[6.x.3.2 MPPSReported (RAD-xx04) 48](#_Toc476134986)

[6.x.4 encoding an event 48](#_Toc476134987)

[6.x.5 Coded Terminologies 51](#_Toc476134988)

[6.x.5.1 Person roles 51](#_Toc476134989)

[6.x.5.1.1 DICOM CID 7450 51](#_Toc476134990)

[6.x.5.1.2 Additional Roles 51](#_Toc476134991)

[6.x.5.2 Machine Roles 52](#_Toc476134992)

[6.x.5.3 Participating Object Roles 52](#_Toc476134993)

[6.x.5.4 Location Roles 53](#_Toc476134994)

[6.x.6 Examples 53](#_Toc476134995)

[6.x.6.1 OrderEntered example 53](#_Toc476134996)

[6.x.6.2 PatientArrived example 55](#_Toc476134997)

[Appendices 57](#_Toc476134998)

[Appendix A – <Appendix A title> 57](#_Toc476134999)

[Appendix B – <Appendix B Title> 57](#_Toc476135000)

[B.1 <Add Title> 57](#_Toc476135001)

[Volume 3 Namespace Additions 57](#_Toc476135002)

[Volume 4 – National Extensions 58](#_Toc476135003)

[4 National Extensions 58](#_Toc476135004)

# Introduction to this Supplement

<Provide a brief overview of the volumes/sections of the Technical Framework that get changed/ added by this supplement. Provide 200 words or less describing this supplement.>

## Open Issues and Questions

|  |  |  |
| --- | --- | --- |
| Number | Issue | |
| 2 | Are there additional events in SWIM that should be required in the profile? Are there other events that should be in the profile (and added to SWIM)? The intention is to harmonize additional events with SWIM.  ***Preliminary list selection done in F2F number one. See section 6.X (volume 3)***  ***(17 Feb) Critical Result discovered added to the list.*** | |
| 5 | Should this profile add the RESTful POST of syslog messages that was proposed for ITI this year. It wasn’t ranked high enough for ITI, and was thought small enough that perhaps a CP was the right approach rather than a supplement. This has two potential uses:   1. A "push" mode for delivering a selection of reports from an Event Record Repository to an Event Consumer. Field experience with security audits is that there is often a need to deliver a subset of the events to a consumer for analysis. Defining this subset in terms of queries can be very complex. It is often much easier for the Repository to select the set of events to be analyzed and send them in a single transaction for analysis. This could also be done by configuring the repository as an event reporter and sending them by means of Syslog. This is very inconvenient for both repository and consumer. A simple RESTful PUT of the set of event reports is preferred. 2. A "bulk push" mode for delivering event reports that have been saved locally for later delivery from an Event Reporter to an Event Record Repository. This may be more efficient or convenient for the Event Reporter, especially for mobile reporters that do not have a native Syslog capability.   The alternative is to set up and then take down syslog connections.  ***Iincluded for PC and decide then. It addresses the issue that Apple and Android both have highly performance optimized (battery, storage, etc.) proprietary logging systems. A mobile device can operate offline or with limited connectivity using the proprietary logging system and then offload the accumulated SOLE log to a repository when the device has good power and connectivity.***  ***We need feedback from ITI as part of PC.*** | |
| 7 | Are events defined so that total event flow remains acceptable for traffic limits of syslog?  Review list in section 6.X and consider what the volume will be like in realistic large environments. | |
| 8 | Should mobile sources be supported in first round using RAD-XX transaction? Eliminates need to establish solution for the mobile devices that will not support syslog.  **YES – If ITI concurs with the approach and is likely to use this for ATNA.** | |
| 12 | (To be answered within IHE after discussion). How do we document compliance and grouping? This is similar to the issues with supporting SOPs, specific CDA formats, etc. The Event Reporter as a generic actor can be documented as complying with the SOLE profile. That is how this supplement is written.  An alternative is to have specific options for various kinds of events. | |
| 15 | In the interest of patient privacy we removed the patientID from the SOLE event, and kept the Exam and Accession number. Will this cause problems with any analysis?  To trace a particular patient you would need to obtain the patient’s exam numbers before patient identity could be derived from the log. How does this affect the ADT phase? Would arrival be logged as "unidentified patient arrived for exams A, B, … ? | |

## Closed Issues

|  |  |  |
| --- | --- | --- |
| Number | Issue | |
| 1 | Should the transport be Syslog, as it is in ATNA?  **Conclusion: Syslog and ATNA compatibility remain the correct choice.**  There are many different transport technologies for event reporting. The IT and other industries have not standardized these well. Problem specific constraints (plus not-invented-here) have driven many variations.  The profile proposes the use of Syslog, consistent with ATNA and allowing re-use of many systems that support ATNA and Syslog.   1. Existing technologies in use for event logging   The following are four examples of event report transport. These are all in widespread use. There are also some important proprietary logging systems in use for mobile devices.  **Syslog**  This protocol has reached standardization, and it’s what DICOM and ATNA use. (So I guess it’s multiply standardized.) It is universally available for laptops and desktop systems. It has limited availability for mobile devices.  Syslog transport has a variety of performance and failure mode characteristics. These are probably acceptable but need to be recognized. They have been acceptable for other audit logs.  Syslog limitations that may matter are:   1. It does not guarantee delivery. TLS connection failures may result in truncated or lost messages. (These are quite rare.) It is the event source responsibility to manage buffering and re-transmission in the event of connection losses. 2. The overhead per message is low, but a connection must be maintained. This is primarily a concern for mobile and battery powered devices that are event reporters. For mobile and battery powered devices a two step approach can be used. The first step is an internal journaling protocol defined and optimized for battery powered and disconnected operation. The second step is a network transmission from the internal log to a syslog server when power and network connectivity is available. Both Apple and Android have switched to a proprietary logging system, and away from Syslog, in part for this reason. 3. Throughput is limited by network speed and buffering. This typically limits events to a few hundred to a few thousand per second. Higher volume event logs from large scale computing are one motivation for FLUME and similar high volume transports.   **Journalctl**  This is the emerging standard in Linux systems for internal journaling. Journalctl implementations include tools to send selections from the journalctl logs to external journaling systems by means of Syslog. There are various motivations for this new journaling system. The primary ones (plus of course not-invented-here) are throughput, memory management, power management, and I/O management. An ordinary workstation will generate thousands of events per second during busy times such as system startup.  They wanted to keep the performance impact low, and a network based system like syslog did not fit.  There is no official standard organization responsible for this logging system. Journalctl is controlled by the Linux Foundation, and is designed to meet their goals:   * High performance integration with the Linux kernel in both large and small configurations. This makes it much less interesting to the Windows or MacOS users. Android uses the Linux kernel, but Android vendors have not adopted Journalctl. The Linux Foundation is less interested in making adaptations to make journalctl easy to deploy in a Windows or Apple environment. * The expected environment is use in the warehouse, rack, desktop, or laptop environments. Power management is a secondary goal. The primary performance goal is dealing with the very high event loads that can occur in the larger systems and enabling high resolution event recording from a large number of internal sources. Another goal is operational independence between the internal journaling with journalctl and external export by network protocols. * Change management decisions are made internally within the Linux Foundation and Linux kernel developers. While this is a highly visible public process, it is not open to other participant feedback. As a result, some changes have been rather controversial or unexpected at times.   Integration of Journalctl with Syslog, and some other protocols, is one of the design goals. Journalctl can be configured to transmit selected journalctl events to Syslog, either as they occur or later in batch mode. Journalctl defines an export format for media as well.  Selection of Syslog by IHE does not interfere with the use of Journalctl by Linux based systems because they can be configured to export from the journalctl system to Syslog.  **No – because this is not likely to be used outside of Linux.**  **Flume**  Flume is a logging system being pushed by the Apache foundation. Its primary target is massive continuous dataflows from many sources into a distributed database such as Hadoop for processing by distributed queries in environments like warehouse scale computing. An example use would be the collection of all the journalctl traffic for 10,000 nodes onto a distributed database of 100 database servers for analysis by 25 query/analysis nodes. It requires the Event Consumer to be a Hadoop server or equivalent. **It’s a Java only library** at the moment, not a basic communication protocol, which requires the Event Reporter to be at least partially written in Java. (The functionality is a mix of Java processing in a standard library and communications elements, which makes it difficult to separate the library from the communications.)  **There is no official standard organization responsible for this logging system.**  **Bitcoin/Blockchain**  It may feel odd to think of Bitcoin as a logging system, but what it records are events, mostly financial events like "A gave 2 bitcoins to B". It is dealing with an environment where there is massive mutual lack of trust, the transaction log must be redundant and survive extensive attacks, and the transaction volume is small.  The performance tradeoffs have prioritized redundancy, survival of the logs in the face of corrupt and malicious nodes, while accepting high computational burdens and significant connectivity requirements on individual nodes. The protocol requires regular connectivity with many other network nodes.  **There is no official standard organization responsible for this logging system.**  **The high computational burden and connectivity requirements make it unattractive.** *The survival of logs in the face of corrupt, defective, and malicious nodes is of value.* The computational cost and connectivity requirements would be a serious problem for mobile use. For other systems it might be feasible, but it is not clear that the risk of corrupt, defective, and malicious nodes is sufficient to justify the cost.  **Android/Ios (phone/tablet use)**  Both Google and Apple have switched from using Syslog to using internal logging methodologies where battery performance can be optimized. Neither chose journalctl. They have both a two step system where the first step is internal logging that is performed using the internal battery optimized method, and then a second step that extracts selected relevant events from that log and sends them to another network node using another method. This other method is typically able to send large blocks of event reports efficiently.  There is no official standard organization responsible for these logging systems. Both are proprietary. Both have interesting capabilities that are closely integrated with the Android or IOS operating system environment. For example, Apple has an internal tagging and data concealing system to manage logs that contain private information and minimize the disclosure of private information to applications that process the logs.  **A bulk event transfer transaction is proposed to deal with this. RAD-XX. This remains a risk issue until ITI confirms that they will use this.**  **MQTT (Oasis)**  MQTT is a lightweight TCP transaction system designed for IoT type applications. From the MQTT web site:  MQTT is a machine-to-machine (M2M)/"Internet of Things" connectivity protocol. It was designed as an extremely lightweight publish/subscribe messaging transport. It is useful for connections with remote locations where a small code footprint is required and/or network bandwidth is at a premium. For example, it has been used in sensors communicating to a broker via satellite link, over occasional dial-up connections with healthcare providers, and in a range of home automation and small device scenarios.  MQTT is standardized as an OASIS standard http://docs.oasis-open.org/mqtt/mqtt/v3.1.1/os/mqtt-v3.1.1-os.pdf | |
| 3 | Should the message be encoding different than the DICOM Audit Schema be used?  **No - The profile assumes use of the DICOM Audit Schema so that ATNA tooling can be re-used**. | |
| 4 | ***Removed, editorial instructions not a comment issue*** | |
| 6 | Are the RESTful Syslog query parameters sufficient? Are additional query parameters needed? They are very coarse query controls. The primary query is to obtain all syslog reports over a time interval from specific sources. The query is against the Syslog message header parameters, and does not query inside the event message.  The role of the server is to maintain the log and provide manageable "chunks" of data to the client. The server does not do detailed filtering. Filters include time range, source, application, message type, and regular expression match of message contents. Final filtering and reporting remain a client task.  The bulk of event selection and analysis is performed within the Event Consumer.  See section 4.1.3 for details of basic query capability. | |
| 9 | We will not capture patient prep times in detail. In general, the time from arrival to procedure room is treated as one activity. Most automatic systems do not collect details within this activity. They cannot distinguish between transport time, time in waiting rooms, and clinical preparation time. | |
| 10 | ***Duplicate of 5, merged into 5*** | |
| 11 | A requirement for NTP on mobile devices may be hard to meet. Should this ATNA requirement be relaxed and rephrased? For cellular network devices it is nearly impossible to meet. Cellular network devices are usually synchronized with the cellular network time base. The GSM standard timebase is not UTC and does not use NTP. It is a good and sufficient time base, but it is 10 to 20 seconds different than UTC. (GSM does not adjust for leap seconds.) For WiFi-only connected devices NTP may be practical.  The grouping with Time Client is to ensure that all the event reports use the same time base. Is the small consistent offset between UTC and GSM time acceptable? Similar issues arise if the mobile device synchronizes with GPS time. It is also stable and consistent, but 10-20 seconds different than UTC.  **Close for RAD, move to ITI. Need to have the same answer for all CT devices.** | |
| 13 | This profile will not cover activities subsequent to completion and delivery of the imaging report. The total list of events in SWIM does extend further, and it could be enhanced with more events for that purpose. That will be the scope of some other profile. | |
| 14 | Should SOLE also use the "IHE+RFC3881" identifier or use a distinct identifier "IHE+SOLE"? One reason to use IHE+SOLE would be to make it easy to do a single query that means "Find all SOLE messages" without parsing the message payloads. The Syslog protocol anticipates that there will be two levels of message identifier that are left to the applications to define. Some commercial Syslog products are configurable to use these levels for filtering and forwarding decisions without needing additional options that parse the contents of the payload. | |
| 16 | This proposes a "IHE+SOLE" EventIDin the Syslog headers, so that combined events can be reported by multiple EventTypeCodes in the MSG-ID. | |
| 17 | Should offset and limit support be mandatory?  **Yes -** | |
|  |  | |

# General Introduction

Update the following appendices to the General Introduction as indicated below. Note that these are not appendices to Volume 1.

Appendix A – Actor Summary Definitions

Add the following actors to the IHE Technical Frameworks General Introduction list of Actors:

<Add any actor definitions for new actors defined specifically for this profile. These will be added to the IHE TF General Introduction list of Actors namespace.>

| Actor | Definition |
| --- | --- |
| Event Reporter | The Event Reporter composes and sends event reports to other actors. |
| Event Consumer | The Event Consumer processes event reports. It may either receive event reports or query for event reports. |
| Event Repository | The Event Repository receives and manages event reports. |

Appendix B – Transaction Summary Definitions

Add the following transactions to the IHE Technical Frameworks General Introduction list of Transactions:

<Add any transaction definitions for new transactions defined specifically for this profile. These will be added to the IHE TF General Introduction list of Transactions namespace.>

| Transaction | Definition |
| --- | --- |
| Transfer Multiple Event Reports | Delivers a payload of many event reports as a single RESTful HTTP Put transaction. |
|  |  |

Glossary

Add the following glossary terms to the IHE Technical Frameworks General Introduction Glossary:

<Any glossary additions associated with the profile draft go here.>

| Glossary Term | Definition |
| --- | --- |
|  |  |
|  |  |

Volume 1 – Profiles

## <*Copyright Licenses>*

<General copyright licenses and permissions are listed in the IHE Technical Frameworks General Introduction. Add information on any standards referenced in the profile that are not already addressed in the permission section.>

Add the following to the IHE Technical Frameworks General Introduction Copyright section:

## <*Domain-specific additions>*

<Some domains have specific sections, added as subsections to Sections 1 or 2, in their Technical Frameworks. These types of additions are allowed as long as they do not adjust the overall numbering scheme which needs to remain consistent across domains. If there are such additions, they should be included here.>

Add to section X …

# X Standardized Operational Log of Events (SOLE) Profile

Efficient businesses use business intelligence tools to manage their business. The application of these tools to manage medical care has been limited in part because the information often resides in several different systems, and there are not standard ways to obtain the information. The SOLE profile defines a way to exchange information about events that can then be collected and displayed using standard methods.

Healthcare providers have a strong desire to increase throughput and efficiency, both to improve the quality and timeliness of care and to control costs. Such process improvement efforts depend on being able to capture workflow events and apply business intelligence tools. Such initiatives face several problems:

* Event information that is to be logged comes from many different systems but there is no easy way to collect and compile the event reports into a single collection
* The different systems recording the particular events being reported may have different understandings of the definition of the event, time point or measurement; the result is:
* Within a single institution, data is non-uniform across systems, degrading the value of the information
* Across institutions, it is hard to compare to evaluate best practices

## X.1 SOLE Actors, Transactions, and Content Modules

This section defines the actors, transactions, and/or content modules in this profile. General definitions of actors are given in the Technical Frameworks General Introduction Appendix A at [http://ihe.net/Technical\_Frameworks](http://ihe.net/Technical_Frameworks/).

Figure X.1-1 shows the actors directly involved in the SOLE Profile and the relevant transactions between them. If needed for context, other actors that may be indirectly involved due to their participation in other related profiles are shown in dotted lines. Actors which have a mandatory grouping are shown in conjoined boxes.

ITI-82: Retrieve Syslog Event

ITI-20: Record Audit Event

RAD-XX: Transfer Multiple Event Reports

ITI-81: Retrieve ATNA Audit Event

Event Reporter

Event Repository

Event Consumer

ITI-20: Record Audit Event

RAD-XX: Transfer Multiple Event Reports

Table X.1-1 lists the transactions for each actor directly involved in the SOLE Profile. To claim compliance with this Profile, an actor shall support all required transactions (labeled "R") and may support the optional transactions (labeled "O").

Table X.1-1: SOLE Profile - Actors and Transactions

| Actors | Transactions | Optionality | Reference |
| --- | --- | --- | --- |
| Event Reporter | ITI-20 Record Audit Event | R | ITI TF-2: 3.20 |
| RAD-XX Transfer Multiple Event Reports | O | 3.Y |
| Event Repository | ITI-20 Record Audit Event | R | ITI TF-2: 3.20 |
| RAD-XX Transfer Multiple Event Reports | R | 3.Y |
| ITI-82 Retrieve Syslog Event | R | ITI TF-2: 3.82 |
| ITI-81 Retrieve ATNA Audit Event | R | ITI TF-2: 3.81 |
| Event Consumer | ITI-20 Record Audit Event | O | ITI TF-2: 3.20 |
| RAD-XX Transfer Multiple Event Reports | O | 3.Y |
| ITI-82 Retrieve Syslog Event | R | ITI TF-2: 3.82 |
| ITI-81 Retrieve ATNA Audit Event | O | ITI TF-2: 3.81 |

Note: The Event Repository is required to implement ITI-20 and RAD-XX as both the Sender and the Receiver.

Table X.1-2: SOLE Profile - Actors and Content Modules

| Actors | Content Modules | Optionality | Reference |
| --- | --- | --- | --- |
| Event Reporter | SOLE Event | R | 6.X SOLE Event Definitions |
| Event Repository | SOLE Event | R | 6.X SOLE Event Definitions |
| Event Consumer | SOLE Event | R | 6.X SOLE Event Definitions |

### X.1.1 Actor Descriptions and Actor Profile Requirements

Most requirements are documented in Transactions (Volume 2) and Content Modules (Volume 3). This section documents any additional requirements on profile’s actors.

#### X.1.1.1 Event Reporter

Event Reporters shall be able to report all events in Table 6.X.2-1 Baseline SOLE Events that occur on the Event Reporter and grouped actors.

The Event Reporter shall be configurable as to the events to be reported, and the location or locations to which they are to be reported.

The Event Reporter shall detect when network connectivity is not available. When network connectivity is not available, event reports shall still be stored locally for later delivery.

The Event Reporter shall be able to manually or automatically transfer the locally stored event reports to an Event Record Repository when network connectivity is made available.

#### X.1.1.2 Event Repository

Event Repository shall be able to store all event reports in Table 6.X.2-1 Baseline SOLE Events. The Event Repository shall be able to remove event reports after a configurable period of time.

The Event Repository shall not generate an error for messages complying with the Syslog RFCs.

Note: This means that a Syslog RFC compliant event report with a malformed SOLE payloads will not generate error.

#### X.1.1.2 Event Consumer

An Event Consumer shall be able to receive any SOLE event report, Table 6.X.2-1 Baseline SOLE Events. It might only process a subset of the SOLE event reports, and might receive and process other types of event reports.

The Event Consumer processes event reports for analysis or display. This profile does not specify the nature of the analysis or display to be performed. An Event Consumer may be a business process analysis system that is being used to improve a clinical workflow. An Event Consumer may be a "dashboard" system that continuously displays the present state of work activities at a location.

## X.2 SOLE Actor Options

Table X.2-1: Standardized Operational Log of Events - Actors and Options

| Actor | Option Name | Reference |
| --- | --- | --- |
| Event Reporter | Multiple Events Option | X.2.1 |
| Event Consumer | Multiple Events Option | X.2.1 |
| Retrieve ATNA Audit Event | X.2.2 |

### X.2.1 Multiple Event Option

An Event Report that supports the Multiple Events Option shall be able to send event reports by means of the Transfer Multiple Event Reports [RAD-XX] transaction. This option does not specify the user interface or system logic used to determine what event reports are to be sent or how the destination is selected or configured.

An Event Consumer that supports the Multiple Events Option shall be able to accept event reports sent by means of the Transfer Multiple Event Reports [RAD-XX] transaction. This actor shall be able to process or store these event reports in the same way as event reports sent by other means.

## X.3 SOLE Required Actor Groupings

It is important that the time base for all of the actors be the same. The Event Reporter and Event Consumer shall be grouped with the Time Client for this reason.

Table X.3-1: Standardized Operational Log of Events - Required Actor Groupings

| SOLE Actor | Actor to be grouped with | Reference | Content Bindings Reference |
| --- | --- | --- | --- |
| Event Reporter | Time Client | ITI TF-1: 2.2.7 | -- |

## X.4 SOLE Overview

### X.4.1 Concepts

The SIIM Workflow Initiative in Medicine (SWIM™) is a Society for Imaging Informatics in Medicine sponsored initiative with initial goals of developing:

1. common definition of the workflow steps within medical imaging departments;
2. Key performance indicators that are defined using these workflow steps;
3. Definitions of the data elements used to capture information about the workflow steps associated with these key performance indicators; and

#### X.4.1.0 Events and extending events.

SOLE has defined a short list of events of interest that take place during imaging by imaging departments. This list can be found in Table 6.X.2-1 Baseline SOLE Events. This list is being coordinated with the RADLEX SWIM.

The RADLEX SWIM initiative has defined and maintains a large list of events that take place in the imaging departments. The full list of events defined by SWIM can be found at <http://siim.org/resource/resmgr/swim/SWIMRadlex1.xlsx>.

The baseline SOLE events will be understood by all SOLE actors.

A deployment may extend the list of events reported for local purposes by extending the baseline SOLE list with reports for other events in the SWIM list. These additional event report types are more likely to be interoperable and understood by other implementations.

New event reports types that are not defined by SWIM are best handled by coordinating with SWIM to get those events incorporated into the SWIM definition.

Local extensions to the event types reported will not generate errors, but are less likely to be interoperably supported.

#### X.4.1.2 Query

The consumers need to query. There are two ways because FHIR and Syslog. Both of the query mechanisms are transactions defined in the ITI Technical Framework. Their capabilities are summarized here.

##### X.4.1.2.1 Retrieve Syslog Message (ITI-82)

The Retrieve Syslog Message (ITI-82) retrieves all syslog messages that match query parameters against the mandatory Syslog header fields. This retrieves all matching syslog messages regardless of the message body format, so it can retrieve any valid Syslog message.

Table 3.82.4.1.2.2-1: Retrieve Syslog Event Keys

| Syslog RFC 5424 element | Retrieve Syslog Event Search Parameter |
| --- | --- |
| PRI | Major syslog category and severity, eg., "103" means security system, error condition. |
| VERSION | Syslog protocol version |
| TIMESTAMP | Time of message (date range query is supported) |
| HOSTNAME | Name of machine that originated the syslog message |
| APP-NAME | Identification for kind of message. IHE has specified "IHE+RFC-3881" for ATNA and "IHE+SOLE" for SOLE event reports. RFC 5424 does not specify the rules for APP-NAME other than specifying that they are application defined strings for use in selecting and filtering. |
| PROCID | Flag to identify logging discontinuities. Typically a process ID for a syslog process. |
| MSG-ID | Typically message type within APP-NAME. RFC 5424 does not specify the rules for MSG-ID other than specifying that they are application defined strings for use in selecting and filtering. SOLE has specified that this will be the SOLE event ID code., e.g., "RID4585". |
| MSG | Message body regular expression match. |

##### X.4.1.2.2 Retrieve ATNA Audit Record (ITI-81)

The ATNA Retrieve Audit Event transaction support searches based on:

* **Patient identifier**: this search parameter allows discovering all of the event reports that occurred related to a specific patient;
* **User identifier**: this search parameter allows discovering all of the actions performed by a specific user
* **Object identifier**: this search parameter allows discovering each event report that occurred related to a specific object (like study, reports, image, etc.).
* **Time frame**: this search parameter allows discovering all of the event reports that occurred during a specific time frame.
* **Event type**: this search parameter allows discovering all of the occurrences of a specific event (like Data Export, Data Import, Query, Authentication, etc.).
* **Application identifier**: this search parameter allows discovering all of the event reports recorded by a specific application or system.
* **Event Outcome Indicator**: this search parameter allows discovering all of the event reports characterized by a specific outcome (Success, Failure, etc.) of the related event.

These retrievals are based on using the FHIR DSTU-2 Audit Record Resource and performing a constrained set of FHIR queries. It returns a set of FHIR resources, per FHIR formatting.

#### X.4.1.3 Filter and Forward

Audit Repositories can provide filtering and forwarding of event reports to other Audit Repositories and to Event Consumers for any purpose; it is not just for dashboards. The Audit Repository is configured with information about the Event Consumers that should receive a selection of event reports. This selection and filtering capability sometimes involves examination of the detailed contents of the event reports.

Filtering and forwarding has been used to create federated reporting systems, and to allow sharing of one audit repository for multiple independent event reporting streams. An Audit Repository may be gathering security and privacy event reports in accordance with IHE ATNA, proprietary database events, facility maintenance events, etc. All of these can use Syslog and the Audit Repository.

The Filtering and Forwarding permit the audit repository to be configured to selectively forward these event reports on to multiple event consumers and audit repositories. Syslog event streams like selected ATNA (e.g., user login) and SOLE event reports might be combined and sent to an integrated activity dashboard, while a separate complete ATNA event report stream is also sent to the security office for use by the security system dashboard.

This profile requires that there be filtering and forwarding capabilities. This profile does not specify feature requirements for the filtering capabilities of an Audit Repository. The specific filtering features are part of the product feature set determined by a vendor.

#### X.4.1.4 RESTful Delivery (mobile)


Mobile platforms have connectivity, battery, and storage considerations that have driven the Google and Apple to create proprietary logging systems that operate locally and require applications to fetch the local logs for transmission by some other means. This profile defines a RESTful delivery mechanism for that purpose.

### X.4.2 Use Cases

The Event Repository receives event reports and stores them. It may be part of a federated network of repositories. It may have analysis and reporting capabilities, but those capabilities are not specified as part of this profile. This profile also does not specify the capacity of an Event Repository, because the variety of deployment needs makes it impractical to set requirements for the event report volume or capacity needed.

#### X.4.2.1 Use Case #1: Track Study Reading Activities

The following scenario shows events from a routine radiology reading activity that follows the IHE Radiology Scheduled Workflow and reports events defined by SOLE. The Actor that sends the event report is an Event Reporter that is grouped with the Actor shown in the diagram.



Order-perform-study.txt

@startuml

participant "Event Repository"

"Order Placer" -> "Event Repository" : "ITI-20: (place order)"

"Order Filler" -> "Event Repository" : "ITI-20: (procedure scheduled)"

"ADT" --> "Event Repository" : "ITI-20 (Patient Arrival)"

"Patient Location\nTracking Supplier" --> "Event Repository" : "ITI-20: (patient in)"

"Patient Location\nTracking Supplier" --> "Event Repository" : "ITI-20: (procedure prepared)"

"Acquisition Modality" --> "Event Repository" : "ITI-20: (procedure started)"

"Acquisition Modality" --> "Event Repository" : "ITI-20: (first image collected)"


"Acquisition Modality" --> "Event Repository" : "ITI-20: (Imaging complete)"

"Patient Location\nTracking Supplier" --> "Event Repository" : "ITI-20: (patient out)"

"Image Manager/Archive" --> "Event Repository" : "ITI-20: (QC Completed)"

"Image Manager/Archive" --> "Event Repository" : "ITI-20: (exam prepared)"

@enduml



Routine-report.txt

@startuml

participant "Event Repository"

"Image Display" --> "Event Repository" : "ITI-20: (user login)"

"Image Archive" --> "Event Repository" : "ITI-20: (exam prepared)"

"Image Display" --> "Event Repository" : "ITI-20: (exam opened)"

"Image Display" --> "Event Repository" : "ITI-20: (query images)"

"Image Archive" --> "Event Repository" : "ITI-20: (query images)"

"Image Archive" --> "Event Repository" : "ITI-20: (retrieve images)"

"Image Display" --> "Event Repository" : "ITI-20: (retrieve images)"

"Image Display" --> "Image Display" : "view images and dictate report"

"Image Display" --> "Event Repository" : "ITI-20: (dictated report)"

"Image Display" --> "Event Repository" : "ITI-20: (reported and signed)"

"Order Filler" --> "Event Repository" : "ITI-20: (final publish)"

"Image Display" --> "Event Repository" : "ITI-20: (final publish)"

"Order Filler" --> "Event Repository" : "ITI-20: (transfer to billing)"

@enduml

As each event takes place, the Actors that are grouped with an Event Reporter that complies with the SOLE profile will use the ITI-20 transaction to convey the event description to the Event Repository.


The Event Repository archives these event reports for a configurable period of time so that subsequent analysis can be performed.

#### X.4.2.2 SOLE Event Analysis

An analyst can use SOLE event reports to study the workflow and operations of a facility. The events have already taken place and are archived in the repository. The analyst selects the appropriate time span to be studied and determines the event reporters that may be reporting relevant information. An analysis workstation acts as an Event Consumer and requests all the data from those data sources during that time period. This will include the SOLE event reports and may include a variety of other event reports.

The analyst uses data import tools to ingest the event reports into an appropriate database or analysis system. For example, this might be a free text indexing database, or it could an object database designed to hold SOLE event reports.

The analyst uses this information to generate the analyses and reports of the workflow based on the event reports and other information.



Syslog-query.txt

@startuml

"BI Analyst" --> "Event Consumer" : "Desired analysis"

"Event Consumer" --> "Event Consumer" : "Determine appropriate time period and data sources"

"Event Consumer" --> "Event Repository" : "ITI-85 Request all data for time period and data sources"

"Event Repository" --> "Event Consumer" : "ITI-85 Large data object response"

"Event Consumer" --> "Event Consumer" : "Ingest into analysis DBMS, e.g.,\nfree text query, xml query, etc."

"Event Consumer" --> "BI Analyst" : "Reports and analyses"

@enduml


#### X.4.2.3 Delayed Event Delivery (mobile)

A mobile device that includes an event reporter can operate without a network connection to the event repository. This is often normal operating mode for battery operated systems.

While operating without a connection, the mobile device accumulates internally formatted event reports for those SOLE events that take place. This internal format uses the proprietary internal logging format provided by the mobile device OS, so that battery use and storage use can be optimized as designed by the device OS vendor.

The mobile device is returned to "home base" where it has a reliable network connection. The mobile device operator instructs it to perform the "end of shift" operations. These operations include the retrieval of the SOLE event reports from the internal log storage, reformatting into the SOLE event report format, and transfer from the mobile device to the event repository.

After successfully transferring the event reports to the Event Repository the mobile device cleans its internal log storage in preparation for future mobile activity.



Mobile-push.txt

@startuml

Participant User

"Event Reporter (mobile)" --> "Event Reporter (mobile)" : "Events sent to Internal system logs"

"User" --> "Event Reporter (mobile)" : "It's now time to deliver sync to server"

"Event Reporter (mobile)" --> "Event Repository" : "[RAD-XX]Transfer Multiple Event Reports"

"Event Reporter (mobile)" --> "Event Reporter (mobile)" : "Internal log clean-up for events sent"

@enduml


Figure X.4.2.2-1: Basic Process Flow in <Profile Acronym> Profile

#### X.4.2.4 Dashboard

A clinic maintains operational awareness for their staff by maintaining a "dashboard" that displays the present status of equipment, waiting room queues, processing queues, reporting queues, etc. This display is regularly updated to reflect the present situation as patients arrive, are imaged, and depart.

This dashboard receives SOLE event reports from the Event Repository as they are received by the Event Repository. The Event Repository filters the event reports that it receives to eliminate reports that should not be displayed, e.g., user login reports and security audit reports, and to eliminate SOLE reports from sources that should not be displayed. The Event Repository filters are also configured and updated to reflect changes to the network but content related filters such as "do not report about this kind of exam" may be beyond the ability of an off the shelf Event Repository. The Event Repository forwarding is configured and updated to reflect changes to the locations of the dashboards.

The dashboard system uses the SOLE event reports to maintain and update the internal status description used generate the visual graphics of the dashboard. The dashboard system has other administrative facilities to correct for network downtime, lost event reports, etc.




Same as above pix-rest-query.txt

#### X.4.1.4 RESTful Delivery (outside analysis)

An outside independent auditor requests the SOLE data be provided for a specific time period, e.g. "March", and set of systems. This is an authorized and approved release of data, but the auditor’s access to the network and facility must be limited. Local policy is that rather than provide the auditors direct access to the facility, a data extract will be provided to the auditor.

The Event Repository has an administrative facility that allows it to perform a RESTful HTTP PUT of the SOLE event reports from a specified time period. The auditor has an Event Consumer that has a minimal HTTP server ability to receive an HTTP PUT. The operators of the Event Repository use RAD-XX to send the requested data to the auditor’s simple Event Consumer.

Subsequent data ingest, analysis, and reporting are performed by the independent auditor on the auditor’s system.



Analysis-push.txt

@startuml

"Independent Auditor" --> "Local Administrator" : "Provide all SOLE data for March to my Event Consumer"

"Local Administrator" --> "Event Repository" : "administrative commands"

"Event Repository" --> "Event Consumer" : "RAD-xxx Large Data Transfer"

"Event Consumer" --> "Event Consumer" : "Ingest into analysis DBMS, e.g.,\nfree text query, xml query, etc."

"Event Consumer" --> "Independent Auditor" : "Reports and analyses"

@enduml

### X.4.3 Contents of SOLE messages

The SOLE event report content is motivated by the reports and informational displays that are to be generated from these reports. The starting point for these event definitions and selection is the work done by the SIIM Workflow initiative, see <http://siim.org/?page=swim> .

Typical uses for these reports include:

- Maintaining a visual dashboard for an imaging department. This dashboard is visible to all staff, and it shows the current utilization status for all of the procedure rooms and the waiting line information for all of the equipment. Reporting backlog and current reporting timelines are presented in summary form. Dashboards like this are used as part of daily management of facilities.

- Business analysts use a selection of event reports to generate a utilization report for equipment. This report indicates how well equipment is being utilized at various times during the day, the statistics of waiting times, the occurrence of problems, etc. The analysts correlate this information with other activity to suggest process improvements and to assess the effectiveness of process improvements.

- A per shift backlog report can be generated at the start of each shift to help assign resources and deal with pending work activities.

- Business analysts use a selection of event reports to generate a statistical analysis of how much time it takes from creating an order to having a finished report. These statistics may consider the nature of the exam, the patient, the shift, the equipment used, etc. The analysts correlate this information with other activity to suggest process improvements and to assess the effectiveness of process improvements.

- Similarly, for an outpatient clinic the statistics regarding time from patient arrival to finished report may be analyzed.

## X.5 SOLE Security Considerations

Local security and privacy service protections and user access controls.

## X.5.1 Security Considerations for Actors

The operational security considerations for a content module are dependent upon the security provisions defined by the grouped actor(s).

### X.5.2 Security Considerations for Event Reports

## X.6 SOLE Cross Profile Considerations

The ITI-20 Record Audit Event transaction may be used extended by other profiles. This may result in a mixing of SOLE event reports with other types of event reports. The Event Repository will need to be prepared for this. The Event Consumer may find event reports defined in other profiles that happen to match a query and that get returned with the other matching reports.

Volume 2 – Transactions

Add section 3.Y

## 3.Y Transfer Multiple Event Reports [RAD-XX]

Transfer Multiple Event Reports [RAD-XX] delivers syslog messages in bulk as a single RESTful transaction. The payload format is the same as for Retrieve Syslog Message [ITI-82].

### 3.Y.1 Scope

This transaction is used to provide a bulk transfer of event records, audit records, and other records in syslog format. It may be for transfer of bulk stored event records to a repository, transfer between repositories, or transfer to analysis systems.

RAD-XX

### 3.Y.2 Actor Roles

The Roles in this transaction are defined in the following table and may be played by the actors shown here:

Table 3.Y.2-1 Actor Roles

|  |  |
| --- | --- |
| **Role:** | Sender:  Sends many event records |
| **Actor(s):** | The following actors may play the role ofSender:  Event Reporter:  Event Repository: |
| **Role:** | Receiver:  Receives and processes many event records |
| **Actor(s):** | The following actors may play the role of Receiver:  Event Repository:  Event Consumer: |

Transaction text specifies behavior for each Role. The behavior of specific Actors may also be specified when it goes beyond that of the general Role.

### 3.Y.3 Referenced Standards

[RFC1945] IETF. May 1996. Hypertext Transfer Protocol Version 1.0 (HTTP/1.0) . http://tools.ietf.org/html/rfc1945 .

[RFC2818] IETF. May 2000. HTTP Over TLS. http://tools.ietf.org/html/rfc2818 .

[RFC5234] IETF. January 2008. Augmented BNF for Syntax Specifications: ABNF. <http://tools.ietf.org/html/rfc5234>

[RFC7230] IETF. June 2014. Hypertext Transfer Protocol (HTTP/1.1): Message Syntax and Routing. http://tools.ietf.org/html/rfc7230 .

[RFC7231] IETF. June 2014. Hypertext Transfer Protocol (HTTP/1.1): Semantics and Content. http://tools.ietf.org/html/rfc7231 .

[RFC7232] IETF. June 2014. Hypertext Transfer Protocol (HTTP/1.1): Conditional Requests. http://tools.ietf.org/html/rfc7232 .

[RFC7233] IETF. June 2014. Hypertext Transfer Protocol (HTTP/1.1): Range Requests. http://tools.ietf.org/html/rfc7233 .

[RFC7234] IETF. June 2014. Hypertext Transfer Protocol (HTTP/1.1): Caching. http://tools.ietf.org/html/rfc7234 .

[RFC7235] IETF. June 2014. Hypertext Transfer Protocol (HTTP/1.1): Authentication. http://tools.ietf.org/html/rfc7235 .

[RFC7236] IETF. June 2014. Initial Hypertext Transfer Protocol (HTTP) Authentication Scheme Registrations. http://tools.ietf.org/ html/rfc7236 .

[RFC7237] IETF. June 2014. Initial Hypertext Transfer Protocol (HTTP) Method Registrations. http://tools.ietf.org/html/rfc7237 .

[RFC7540] IETF. May 2015. Hypertext Transfer Protocol Version 2 (HTTP/2). http://tools.ietf.org/html/rfc7540 .

### 3.Y.4 Interaction Diagram

Sender

HTTP POST Request

Message 1

HTTP POST Response

Message 2

Receiver

Figure 3.Y.4-1 Interaction Diagram

#### 3.Y.4.1 HTTP POST Request

The Sender sends a collection of syslog messages to the Receiver.

The Receiver shall support handling of such messages from more than one Sender. The Sender shall support making requests for more than Receiver.

##### 3.Y.4.1.1 Trigger Events

The Sender initiates this transaction whenever it determines that it has a sufficient collection of syslog reports. For example, a mobile device that uses an internal proprietary event log while operating offline could initiate this transaction when it returns to a home location with good power and network connections.

##### 3.Y.4.1.2 Message Semantics

The message is an HTTP POST request. The Sender is the user-agent. The Receiver is the origin server. The request syntax is:

POST SP /{bulk-syslog-events} SP version CRLF

Content-Type: dicom-media-type CRLF

\*(header-field CRLF)

CRLF

payload

###### 3.Y.4.1.2.1 Resource

The target URL shall be <scheme>://<authority>/<path>/bulk-syslog-events.

###### 3.Y.4.1.2.2 Query Parameters

There are no Query Parameters.

###### 3.Y.4.1.2.3 Request Header Fields

Table 3.Y.4.2.2.3-1: Request Header Fields

|  |  |  |
| --- | --- | --- |
| Header Field | Value | Requirements |
| Content-Type | **application/json** | Required |
| Accept | **application/json** | Required |
| Content-Length | uint | Shall be present if no transfer coding has been applied. Shall be absent otherwise. |
| Transfer-Encoding | encoding | Shall be present if a transfer coding has been applied. Shall be absent otherwise. |

###### 3.Y.4.1.2.4 Request Payload

The request payload shall be present and shall contain all the JSON encoded bulk event records.

The syslog messages shall be encoded as described in ITI-TF2c: Table 3.82.4.2.2-1. See ITI Supplement Add RESTful Query to ATNA.



###### 3.Y.4.1.2.4.1 Example JSON encoding

{ "Events": [

{

"Pri" : "103",

"Version": "1",

"Timestamp": "2015-03-17T00:05"

"Hostname": "nemo.frodo.org"

"App-name": "su"

"Procid": "1234"

"Msg-id": "su-fail"

"Msg": "su root failed for sauron"

},

{

"Pri": "101",

"Version": "1",

"Timestamp": "2015-03-17T00:06"

"Hostname": "nemo.frodo.org"

"App-name": "su"

"Procid": "1234"

"Msg-id": "su-start"

"Msg": "su root succeeded for gandalf"

},

{

"Pri": "101",

"Version": "1",

"Timestamp": "2015-03-17T00:15"

"Hostname": "nemo.frodo.org"

"App-name": "su"

"Procid": "1234"

"Msg-id": "su-stop"

"Msg": "su root completed for gandalf"

}

{

"Pri": "tbd",

"Version": "1",

"Timestamp": "2015-03-17T00:15"

"Hostname": "nemo.frodo.org"

"App-name": "IHE+SOLE"

"Procid": "1234"

"Msg-id": "RID12345"

"Msg": "contents of SOLE message"

}

]

}

##### 3.Y.4.1.3 Expected Actions

The Receiver stores the representations contained in the request payload for further actor specific processing and returns a response.

#### 3.y.4.2 HTTP Post Response

##### 3.y.4.2.1 Trigger Events

Receiver receives an HTTP Post Request.

##### 3.Y.4.2.2 Message Semantics

The message is an HTTP Post Response. The Sender is the user-agent. The Receiver is the origin-server.

The response shall have the following syntax:

version SP status-code SP reason-phrase CRLF

\*(header-field CRLF)

CRLF

[Status Report]

###### 3.Y.4.2.2.1 Status Codes

The response shall have an appropriate status code. Table 3.Y.4.2.2.1-1 contains the most common status codes for this transaction.

Table 3.Y.4.2.2.1-1: Common Status Codes

|  |  |
| --- | --- |
| Status Code | Description |
| 200 (OK) | Indicates that the origin server successfully stored or created at least one of the representations contained in the request payload and is returning a response payload. |
| 201 (Created) | Indicates that the origin server successfully created at least one of the representations contained in the request payload and may be returning a response payload. |
| 202 (Accepted) | Indicates that the origin server successfully validated the request message, but has not yet stored or created the representations in the request payload. The origin server may or may not have validated the payload.  The user agent can use a Query or Retrieve transaction later to determine if the request has completed. |
| 204 (No Content) | Indicates that the origin server successfully stored all the representations contained in the request payload without any modifications and is not returning a response payload. |
| 400 (Bad Request) | Indicates that the origin server did not store any of the representations contained in the request payload because of errors in the request message. For example, an invalid encoding. |
| 409 (Conflict) | Indicates that the request could not be completed due to a conflict with the current state of the target resource. |

###### 3.Y.4.2.2.2 Response Header Fields

Table 3.Y.4.2.2.2-1: Response Header Fields

|  |  |  |
| --- | --- | --- |
| Header Field | Value | Requirements |
| Content-Type | application/json | Required |
| Content-Length | uint | Shall be present if no transfer coding has been applied. Shall be absent otherwise. |
| Transfer-Encoding | encoding | Shall be present if a transfer coding has been applied. Shall be absent otherwise. |

###### 3.Y.4.2.2.3 Response Payload

If the origin server failed to store or modified any representations in the request payload, the response payload shall contain a Status Report describing any additions, modifications, or deletions to the stored representations. The Status Report may also describe any warnings or other useful information.

##### 3.Y.4.2.3 Expected Actions

The sender may process a Status Report if present.

### 3.Y.5 Security Considerations

<Description of the transaction specific security consideration; such as use of security profiles.>

#### 3.Y.5.1 Security Audit Considerations

<This section should identify any specific ATNA security audit event that is associated with this transaction and requirements on the encoding of that audit event. >

##### 3.Y.5.1.(z) <Actor> Specific Security Considerations

<This section should specify any specific security considerations on an Actor by Actor basis.>

Appendices

Appendix A – <Appendix A Title>

Appendix A text goes here.

A.1 <Add Title>

Appendix A.1 text goes here

Appendix B – <Appendix B Title>

Appendix B text goes here.

B.1 <Add Title>

Appendix B.1 text goes here.

Volume 2 Namespace Additions

Add the following terms to the IHE General Introduction Appendix G:

<Please explicitly identify all new OIDs, UIDs, URNs, etc., defined specifically for this profile. These will be added to the IHE TF General Introduction namespace appendix when it becomes available. These items should be collected from the sections above, and listed here as additions when this document is published for Trial Implementation. This section will be deleted prior to inclusion into the Technical Framework as Final Text, but should be present for publication of Public Comment and Trial Implementation.>

Volume 3 – Content Modules

# 5 Namespaces and Vocabularies

Add to section 5 Namespaces and Vocabularies

<Note that the code systems already defined in the Technical Framework of this domain may (but not required) be replicated here just to aid in the supplement review as a standalone document. Also note that the Section 5 table numbers and names are already defined in the TF Volume 3.>

| codeSystem | codeSystemName | Description |
| --- | --- | --- |
| <oid or uid> | <code system name> | <short description or pointer to more detailed description> |
| <oid or uid> | <code system name> | <short description or pointer to more detailed description> |
| <oid or uid> | <code system name> | <short description or pointer to more detailed description> |











# 6 Content Modules

## 6.X SOLE Event Definitions

The events described in SWIM <http://siim.org/resource/resmgr/swim/SWIMRadlex1.xlsx> are reported as SOLE event reports. The semantic content to be included in the report is defined in event semantics, and the encoding of those semantics defined in event report encoding.

### 6.X.1 SWIM and SOLE Event selection

The SWIM lexicon was developed as a comprehensive list of all events that might occur in an imaging department. However, many of these are not easily captured by systems that exist today, or the effort to report the events would far outweigh the perceived value. The events were also selected to cover a broad range of event types in order to assure broad utility and to assure the methods selected in the profile would not impede future inclusion of the rest of the SWIM lexicon. Finally, events that are considered highly valuable based on current business intelligence systems were also identified and included in the list.

### 6.X.2 Event Semantics

The baseline events are selected from the RADLEX SWIM list, <https://github.com/ImagingInformatics/SWIM-Events>. This list may be extended by local policy with other events from the RADLEX SWIM list, or by locally defined events.

Table 6.X.2-1 Baseline SOLE Events

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Name | RADLEX ID | Definition: "The time when:" | Actor | Data Field for time (DICOM/HL7) |
| **OrderEntered** | RID45813 | The exam order is entered into the order placer system | **Order Place**r, | **HL7-ORM-ORC segment field no 9** |
| AppointmentTimeScheduled | RID45814 | The scheduler set the ApptTime for the exam | ??? |  |
| **PatientArrived** | RID45825 | The patient checks-in at an identified desk for the exams | **ADT** |  |
| RoomAssigned | RID45934 | The room (?or performing resource?) is assigned to a patient and procedure | ??? |  |
| PatientIn | RID45897 | the patient enters the procedure room | Modality? |  |
| PatientOut | RID45899 | the patient leaves the procedure room | Modality? |  |
| **FirstImageCollected** | RID46000 | the imaging device begins to collect data | **Modality** |  |
| **ImagingComplete** | RID45835 | All images are acquired and reconstructed (including routine additional reconstructions/reformations done on the imaging device) on the imaging device. | **Modality** | MPPS |
| **StudyPrepared** | RID45914 | All steps required for reporting completed (images acquired, transmitted to reporting device, post-processing done) e.g. Exam put onto reading or QC worklist. | RIS? Workstation? |  |
| QCCompleted | RID28816 | Delete unacceptable images, adjust W/L, confirm correct patient, etc. Typically done by a technologist | RIS? Workstation? |  |
| Crit Notification | RID45854 | A category 1 (ACR Definition) finding is communicated to a physician caring for the patient. | FUNC xyz?? |  |
| ReportDictated | RID45859 | Physician reviews image and renders a report in electronic audio format | **RIS** |  |
| **ReportApproved** | RID45924 | Final text form report is approved(signed) | **RIS** |  |
| ReportPublished | RID45865 | Final report and exam is sent to ordering physician (EMR confirmation of receipt) | **EMR,** *RIS?* |  |
| ExamArchiveCommit | RID4580 | Exam is transferred to an external archive or VNA | **Image Archive** |  |
| ExamTransferToBilling | RID45836 | Information on exam has been transferred to billing system | EMR? |  |
| PatientMerged | RID45898 | Exams of a patient with 2 IDs are merged to 1 of the IDs | **EMR** |  |
| ExamReassigned | RID45863 | Change the patient ID for an exam to a different patient ID (e.g., trauma patient ID reconciliation) | **EMR** |  |
| ExamRemoved | RID45856 | The exam and any associated images are deleted (This needs clarification, deleted or just not returned on routine queries? See Exam Deleted) | **EMR** |  |
| ExamOpenedForReporting | RID45893 | The examination is opened on viewing application by the radiologist for reporting | *RIS? PACS?* |  |
| ExamCancelled | RID45862 | Time exam is cancelled (either before or during) the exam | **RIS** |  |
| ExamExceptionDetected | **TBD** | The exam has an exception condition that must be administratively resolved. | **RIS** |  |
| ExamExceptionResolved | **TBD** | The exam exception condition has been administratively resolved. | **RIS** |  |
| ExamDeleted | **TBD** | The exam has been deleted from an application perspective. (This needs clarification, deleted or just not returned on routine queries? See Exam Removed) | **RIS** |  |
| MPPS Message |  | MPPS Reported,  Contents of MPPS |  |  |

#### 6.X.3 Event Reports.

All SOLE event reports shall be reported as APP-NAME of "IHE+SOLE" and MSG-ID of the Event Code, e.g. "RID45859".

#### 6.x.3.1 OrderEntered (RID45813)

The exam order is entered into the order placer system.

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Orderplacer (M)(1) | Patient (M)(1) |
| Ordering Physician (O) (1) | Exam(M)(1..n) |
|  | Accession Number(O)(1..n) |

#### 6.x.3.2 AppointmentTimeScheduled (RID45814)

The scheduler sets the ApptTime for the exam

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Dss/Orderfiller (O)(0..1) | Patient (M)(1) |
| Human Scheduler (O) (0..1) | Exam(M)(1..n) |
| Other automated scheduler (O) (0..1) | Accession Number(O)(0..n) |

At least one Active Participant shall be recorded.

#### 6.x.3.3 PatientArrived (RID45825)

The patient checks-in at an identified desk for the exams

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Admitting Staff(O) (0..1) | Patient (M) (0..1) |
| Automatic Check In System (0..1) | Arrival Location(O) (0..1) |
|  |  |

#### 6.x.3.4 RoomAssigned (RID45934)

The room (?or performing resource?) is assigned to a patient and procedure

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Dss/Orderfiller (o) (0..1) | Patient (M) (1) |
| Human Scheduler (O) (0..1) | Exam(M) (1..n) |
| Other automated scheduler (O) (0..1) | Fixed Location(O) (0..1) |
|  | Resource(O) (0..1) |

At least one Active Participant shall be recorded.

#### 6.x.3.5 PatientIn (RID45897)

The patient enters the procedure room

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Patient (M) (1) | Fixed Location (O) (0..1) |
|  | Mobile Location (O) (0..1) |
|  |  |

#### 6.x.3.6 PatientOut (RID45899)

The patient leaves the procedure room

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Patient (M) (1) | Fixed Location (O) (0..1) |
|  | Mobile Location (O) (0..1) |
|  |  |

#### 6.x.3.7 FirstImageCollected (RID46000)

The imaging device begins to collect data

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Modality (M) (1) | Patient (M) (1) |
|  | Exam(M) (1) |
|  |  |

#### 6.x.3.8 ImagingComplete (RID45835)

All images are acquired and reconstructed (including routine additional reconstructions/reformations done on the imaging device) on the imaging device

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Modality (M) (1) | Patient (M) (1) |
|  | Exam(M) (1..n) |
|  |  |

#### 6.x.3.9 StudyPrepared (RID45914)

All steps required for reporting completed (images acquired, transmitted to reporting device, post-processing done) e.g. Exam put onto reading or QC worklist.

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Technologist(M) (1) | Patient (M) (1) |
|  | Exam(M) (1) |
|  |  |

#### 6.x.3.10 QCCompleted (RID28816)

Delete unacceptable images, adjust W/L, confirm correct patient, etc. Typically done by a technologist

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Technologist(M) (1) | Patient (M) (1) |
|  | Exam(M)(1) |
|  |  |

#### 6.x.3.11 ReportDictated (RID45859)

Physician reviews image and renders a report in electronic audio format

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Physician (M) (1) | Patient (M) (1) |
|  | Exam(M) (1) |
|  | Fixed Location (0..1) |
|  | Mobile Location (0..1) |

#### 6.x.3.12 ReportApproved (RID45924)

Final text form report is approved(signed)

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Physician (M) (1) | Patient (M) (1) |
|  | Exam(M) (1) |
|  |  |

#### 6.x.3.13 ReportPublished (RID45865)

Final report and exam is sent to ordering physician (EMR confirmation of receipt)

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| PACS/RIS(M) (1) | Patient (M) (1) |
|  | Exam(M) (1) |
|  |  |

#### 6.x.3.14 ExamArchiveCommit (RID4580)

Exam is transferred to an external archive or VNA

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Archive(M) (1) | Patient (M) (1) |
|  | Exam(M) (1..n) |
|  | Order(O) (1..n) |

#### 6.x.3.15 ExamTransferToBilling (RID45836)

Information on exam has been transferred to billing system

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| EMR (M) (1) | Patient (M) (1) |
|  | Order(M) (1) |
|  |  |

#### 6.x.3.16 PatientMerged (RID45898)

Exams of a patient with 2IDs are merged to 1 of the IDs

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| PACS/RIS/EMR (M) (1) | Patient (M) (2..n) |
|  | Exam(M) (1..n) |
|  |  |

#### 6.x.3.17 ExamReassigned (RID45863)

Change the patient ID for an exam to a different patient ID (e.g., trauma patient ID reconciliation)

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| PACS/RIS/EMR (M) (1) | Patient (M) (2) |
|  | Exam(M) (1..n) |
|  |  |

#### 6.x.3.18 ExamRemoved (RID45856)

The exam and any associated images are deleted (This needs clarification, deleted or just not returned on routine queries? See Exam Deleted)

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| PACS/RIS/EMR (M) (1) | Patient (M) (1) |
|  | Exam(M) (1..n) |
|  |  |

#### 6.x.3.19 ExamOpenedFoReporting (RID45893)

The examination is opened on viewing application by the radiologist for reporting

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| PACS/RIS (M) (1) | Patient (M) (1) |
| Physician (O) (0..1) | Exam(M) (1) |
|  |  |

#### 6.x.3.20 ExamCancelled (RID45862)

Time exam is cancelled (either before or during) the exam

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| PACS/RIS/EMR (M) (1) | Patient (M) (1) |
|  | Exam(M) (1..n) |
|  |  |

#### 6.x.3.21 ExamExceptionDetected (RAD-xx01)

The exam has an exception condition that must be administratively resolved.

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Detecting System (M) (1) | Patient (M) (1) |
|  | Exam(M) (1) |
|  |  |

#### 6.x.3.22 ExamExceptionResolved (RAD-xx02)

The exam exception condition has been administratively resolved.

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Resolving System (M) (1) | Patient (M) (1) |
|  | Exam(M) (1) |
|  |  |

#### 6.x.3.23 ExamDeleted (RAD-xx03)

The exam has been deleted from an application perspective. (This needs clarification, deleted or just not returned on routine queries? See Exam Removed)

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| PACS/RIS/EMR (M) (1) | Patient (M) (1) |
|  | Exam(M) (1) |
|  |  |

#### 6.x.3.23 CritNotification (RID45854)

A category 1 (minutes) finding (ACR definition, see Actionable Findings and the Role of IT Support: Report of the ACR Actionable Reporting Work Group <http://dx.doi.org/10.1016/j.jacr.2013.12.016>) is communicated to a physician taking care of patient.

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Physician-Sending(M) (1) | Patient (M) (1) |
| Physician-Receiving(M) (1) | Exam(M) (1) |
|  |  |

#### 6.x.3.2 MPPSReported (RAD-xx04)

The modality performing the exam has reported that the procedure is complete.

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Modality(M)(1) | MPPS Contents (1) |
|  |  |
|  |  |

### 6.x.4 encoding an event

Events shall be encoded in accordance with the schema in DICOM PS3.15 Section A.5. The tables associated with each event indicate the active participants and participating objects for each event. These may be mandatory or optional, and may have restrictions on the multiplicity. They shall all comply with the encoding restrictions shown in the table below. There are tables for machines and other conceptual objects (e.g., exam), and tables for humans.

The Event description

|  |  |  |  |
| --- | --- | --- | --- |
|  | Field Name | Opt | Value Constraints |
| **Event**  AuditMessage/ EventIdentification | EventID | M | **EV(SOLE, IHE, "Radiology Operational Event")** |
| EventActionCode | M | **E** |
| *EventDateTime* | *M* | *not specialized* |
| *EventOutcomeIndicator* | *M* | *not specialized* |
| EventTypeCode | M | **DCID(Table 6.X.2-1 Baseline SOLE Events), Multiple EventTypeCodes are permitted.** |
| Audit Source (0..1) | | | |
| Machine Active Participant (0..n) | | | |
| Human Active Participant (0..n) | | | |
| Human Passive Participant, e.g., Patient (0..n) | | | |
| Machine Passive Participant, e.g., Study (0..n) | | | |

Where:

Audit should be included. It is optional in the schema, but it is often useful to specify the system that originally detected and reported the event.

|  |  |  |  |
| --- | --- | --- | --- |
| **Audit Source**  AuditMessage/ AuditSourceIdentification | *AuditSourceID* | *U* | *not specialized* |
| *AuditEnterpriseSiteID* | *U* | *not specialized* |
| *AuditSourceTypeCode* | *U* | *not specialized* |

Machines, software, applications, etc. that actively participate in the event, e.g., Modality or Image Archive shall all comply with the following requirement.

|  |  |  |  |
| --- | --- | --- | --- |
| **Machine Active Participant**  AuditMessage/ ActiveParticipant | UserID | M | One identity of the machine participant, e.g., process ID, AE title, etc. |
| AlternativeUserID | U | A second identity of the machine participant, e.g., process ID, AE title, etc. |
| *UserName* | *U* | *not specialized* |
| *UserIsRequestor* | *U* | *not specialized* |
| RoleIDCode | M | ***See section 6.x.5.2*** |
| NetworkAccessPointTypeCode | M | "1" for machine (DNS) name, "2" for IP address |
| NetworkAccessPointID | M | The machine name or IP address. |

People that actively participate in the event, e.g., Radiologist or Technologist. Additional fields extending DICOM Schema are defined per event. The DICOM CID 7450 provides role codes for the various participation roles. Note that a participation role is different than an organizational role. A surgeon refers a patient for an exam would have an organizational role of surgeon, but a participation role in that exam of referring physician.

|  |  |  |  |
| --- | --- | --- | --- |
| **Human Active Participant**  AuditMessage/ ActiveParticipant | UserID | M | One identity of the human that participated in the transaction, e.g., Employee Number. |
| *AlternativeUserID* | *U* | A second identity of the human that participated in the transaction, e.g., NPI(US). |
| *UserName* | *U* | *not specialized* |
| *UserIsRequestor* | *U* | *not specialized* |
| RoleIDCode | M | **See section 6.x.5.1** |
| Department |  | *See section 6.x.5.5 Departments* |
| Shift |  | *????* |
| NetworkAccessPointTypeCode | *U* | *not specialized* |
| NetworkAccessPointID | *U* | *not specialized* |

System objects, e.g., "exam" that are the objects of actions are described as follows. The object roles are selected from DICOM CID 7445 and from Table 5.1.3.

|  |  |  |  |
| --- | --- | --- | --- |
| **System Object**  (AuditMessage/ ParticipantObjectIdentification) | ParticipantObjectTypeCode | M | "2" (system object) |
| ParticipantObjectTypeCodeRole | M | **See section 6.x.5.3, participating object roles**  **Need table: e.g. Exam, appointment, study, report** |
| *ParticipantObjectDataLifeCycle* | *U* | *not specialized* |
| *ParticipantObjectIDTypeCode* | *M* | *not specialized* |
| *ParticipantObjectSensitivity* | *U* | *not specialized* |
| ParticipantObjectID | *M* | *not specialized* |
| *ParticipantObjectName* | *U* | *not specialized* |
| *ParticipantObjectQuery* | *U* | *not specialized* |
| ParticipantObjectDetail | M | **TBD** |

Location objects, e.g., "Room 101" that are the objects of actions are described as follows. The object roles are selected as specified in Section 6.x.5.3.

|  |  |  |  |
| --- | --- | --- | --- |
| **Participating Location Object**  (AuditMessage/ ParticipantObjectIdentification) | ParticipantObjectTypeCode | M | "2" (system object) |
| ParticipantObjectTypeCodeRole | M | **See section 6.x.5.4, location roles** |
| *ParticipantObjectDataLifeCycle* | *U* | *not specialized* |
| *ParticipantObjectIDTypeCode* | *M* | *not specialized* |
| *ParticipantObjectSensitivity* | *U* | *not specialized* |
| ParticipantObjectID | *M* | *not specialized* |
| *ParticipantObjectName* | *U* | *not specialized* |
| *ParticipantObjectQuery* | *U* | *not specialized* |
| ParticipantObjectDetail | M | **TBD** |

Human objects, e.g., patient. Additional fields extending DICOM Schema and ParticipantObjectDetail requirements are defined per event.

|  |  |  |  |
| --- | --- | --- | --- |
| **Human Participating Object**  (AuditMessage/ ParticipantObjectIdentification) | ParticipantObjectTypeCode | M | "1" (person) |
| ParticipantObjectTypeCodeRole | M | **See section 6.x.5.1. This is usually patient** |
| *ParticipantObjectDataLifeCycle* | *U* | *not specialized* |
| *ParticipantObjectSensitivity* | *U* | *not specialized* |
| ParticipantObjectID | M | the human ID in HL7 CX format. |
| *ParticipantObjectIDTypeCode* | *M* | *not specialized* |
| *ParticipantObjectName* | *U* | *not specialized* |
| *ParticipantObjectQuery* | *U* | *not specialized* |
| ParticipantObjectDetail | M | **TBD** |

### 6.x.5 Coded Terminologies

#### 6.x.5.1 Person roles

##### 6.x.5.1.1 DICOM CID 7450

The DICOM CID 7450 has a list of potential person roles. It include further lists of family member and organizational roles. These may be used as the role for a person in a SOLE message. It is reproduced here for convenience.

**Table CID 7450. Person Roles**

| **Coding Scheme Designator** | **Code Value** | **Code Meaning** | **SNOMED-CT Concept ID** | **UMLS Concept Unique ID** |
| --- | --- | --- | --- | --- |
| DCM | [121025](#DCM_121025) | Patient |  |  |
| SRT | [J-00552](http://browser.ihtsdotools.org/?perspective=full&conceptId1=223366009) | Healthcare professional | [223366009](http://browser.ihtsdotools.org/?perspective=full&conceptId1=223366009) | [C1704312](https://uts.nlm.nih.gov/metathesaurus.html?cui=C1704312) |
| SRT | [S-11090](http://browser.ihtsdotools.org/?perspective=full&conceptId1=113163005) | Friend | [113163005](http://browser.ihtsdotools.org/?perspective=full&conceptId1=113163005) | [C0079382](https://uts.nlm.nih.gov/metathesaurus.html?cui=C0079382) |
| *Include [CID 7451 "Family Member"](#sect_CID_7451)* <http://dicom.nema.org/medical/dicom/current/output/chtml/part16/sect_CID_7451.html> | | |  |  |
| *Include [CID 7452 "Organizational Roles"](#sect_CID_7452) <http://dicom.nema.org/medical/dicom/current/output/chtml/part16/sect_CID_7452.html>* | | |  |  |

##### 6.x.5.1.2 Additional Roles

The following additional roles may be used in SOLE messages.

| **Coding Scheme Designator** | **Code Value** | **Code Meaning** | **SNOMED-CT Concept ID** | **UMLS Concept Unique ID** |
| --- | --- | --- | --- | --- |
|  |  |  |  |  |
|  |  |  |  |  |
|  |  |  |  |  |

#### 6.x.5.2 Machine Roles

The following machine roles should be used for devices that are active participants in an event. If this list lacks a suitable code, a locally defined code may be used.

**Table Device Participating Roles**

| **Coding Scheme Designator** | **Code Value** | **Code Meaning** |
| --- | --- | --- |
| DCM | [113859](#DCM_113859) | Irradiating Device |
| DCM | [121097](#DCM_121097) | Recording |
| DCM | [113942](#DCM_113942) | X-Ray Reading Device |
|  |  |  |
|  |  |  |

#### 6.x.5.3 Participating Object Roles

The following participating object roles should be used for objects that participate in an event. If this list lacks a suitable code, a locally defined code may be used.

**Table   Participating Object Roles**

| **Coding Scheme Designator** | **Code Value** | **Code Meaning** |
| --- | --- | --- |
|  |  | Examination |
|  |  | Appointment |
|  |  | Study |
|  |  | Report |
|  |  |  |

#### 6.x.5.4 Location Roles

The following machine roles should be used for locations that participate in an event. If this list lacks a suitable code, a locally defined code may be used.

**Table Location Roles**

| **Coding Scheme Designator** | **Code Value** | **Code Meaning** |
| --- | --- | --- |
|  |  | Room Identifier |
|  |  | Mobile Facility Identifier |
|  |  |  |
|  |  |  |
|  |  |  |

### 6.x.6 Examples

#### 6.x.6.1 OrderEntered example

An example of a SOLE event report for an OrderEntered is provided below. In this example, Dr. Smith has order an exam for patient Jones. The Exam was assigned the number "ex-1245" and accession number "ac-9383" by the order placer system.

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| Orderplacer (M)(1) | Patient (M)(1) |
| Ordering Physician (O) (1) | Exam(M)(1..n) |
|  | Accession Number(O)(1..n) |

|  |  |  |  |
| --- | --- | --- | --- |
|  | Field Name | Opt | Value Constraints |
| **Event**  AuditMessage/ EventIdentification | EventID | M | **EV(SOLE, IHE, "Radiology Operational Event")** |
| EventActionCode | M | **E** |
| *EventDateTime* | *M* | *2015-03-17T00:15* |
| *EventOutcomeIndicator* | *M* | *"0"* |
| EventTypeCode | M | **EV(RID45813, RADLEX, "OrderEntered")** |

|  |  |  |  |
| --- | --- | --- | --- |
| **Audit Source**  AuditMessage/ AuditSourceIdentification | *AuditSourceID* | *U* | *Orderplacer-2* |
| *AuditEnterpriseSiteID* | *U* |  |
| *AuditSourceTypeCode* | *U* | *"4"* |

|  |  |  |  |
| --- | --- | --- | --- |
| **Machine Active Participant**  **AuditMessage/ ActiveParticipant** | UserID | M | <order-placer-process-ID> |
| AlternativeUserID | U |  |
| *UserName* | *U* |  |
| *UserIsRequestor* | *U* |  |
| RoleIDCode | M | ***EV(TBD, IHE-RAD, "Order Placer")*** |
| NetworkAccessPointTypeCode | M | 1 |
| NetworkAccessPointID | M | Placer2.hospital.org |
| **Human Active Participant**  AuditMessage/ ActiveParticipant | UserID | M | PID-horn1234123 |
| *AlternativeUserID* | *U* |  |
| *UserName* | *U* |  |
| *UserIsRequestor* | *U* |  |
| RoleIDCode | M | **EV( 121025, DCM, "Patient")** |
| Department | *U* | *"Internal Medicine"* |
| Shift | *U* |  |
| NetworkAccessPointTypeCode | *U* |  |
| NetworkAccessPointID | *U* |  |

|  |  |  |  |
| --- | --- | --- | --- |
| **System Object**  (AuditMessage/ ParticipantObjectIdentification) | ParticipantObjectTypeCode | M | "2" (system object) |
| ParticipantObjectTypeCodeRole | M | **EV( examcode, ???, "Exam ID")** |
| *ParticipantObjectDataLifeCycle* | *U* | *not specialized* |
| *ParticipantObjectIDTypeCode* | *M* | **EV( examcode, ???, "Exam ID")** |
| *ParticipantObjectSensitivity* | *U* |  |
| ParticipantObjectID | *M* | *"ex-1245"* |
| *ParticipantObjectName* | *U* |  |
| *ParticipantObjectQuery* | *U* |  |
| ParticipantObjectDetail | U |  |
| **System Object**  (AuditMessage/ ParticipantObjectIdentification) | ParticipantObjectTypeCode | M | "2" (system object) |
| ParticipantObjectTypeCodeRole | M | **EV( examcode, ???, "Accession Number")** |
| *ParticipantObjectDataLifeCycle* | *U* |  |
| *ParticipantObjectIDTypeCode* | *M* | *EV( examcode, ???, "Accession Number")* |
| *ParticipantObjectSensitivity* | *U* |  |
| ParticipantObjectID | *M* | *"ac-9383"* |
| *ParticipantObjectName* | *U* |  |
| *ParticipantObjectQuery* | *U* |  |
| ParticipantObjectDetail | U |  |

#### 6.x.6.2 PatientArrived example

An example of a SOLE event report for a Patient Arrival event is provided below. The example will use an patient who arrives and uses and automatic checkin system at the radiology clinic entrance.

|  |  |
| --- | --- |
| Active Participants (Subject) | Participating Objects (Object) |
| ~~Admitting Staff(O) (0..1)~~ |  |
| Automatic Check In System (0..1) | Arrival Location(O) (0..1) |
| Patient (M) (0..1) |  |

|  |  |  |  |
| --- | --- | --- | --- |
|  | Field Name | Opt | Value Constraints |
| **Event**  AuditMessage/ EventIdentification | EventID | M | **EV(SOLE, IHE, "Radiology Operational Event")** |
| EventActionCode | M | **E** |
| *EventDateTime* | *M* | *2015-03-17T00:15* |
| *EventOutcomeIndicator* | *M* | *0* |
| EventTypeCode | M | **EV( RID45825, RADLEX, "PatientArrived"** |

|  |  |  |  |
| --- | --- | --- | --- |
| **Audit Source**  AuditMessage/ AuditSourceIdentification | *AuditSourceID* | *U* | *AUTOADMIT* |
|  |  |  |
| *AuditSourceTypeCode* | *U* | *"4"* |

|  |  |  |  |
| --- | --- | --- | --- |
| **Machine Active Participant**  AuditMessage/ ActiveParticipant | UserID | M | <automatic-admitting-process-ID> |
| AlternativeUserID | U |  |
| *UserName* | *U* |  |
| *UserIsRequestor* | *U* |  |
| RoleIDCode | M | ***See section 6.x.5.2*** |
| NetworkAccessPointTypeCode | M | "1" |
| NetworkAccessPointID | M | "Autoadmit7.hospital.org" |
| **Human Active Participant**  AuditMessage/ ActiveParticipant | UserID | M | PID-horn1234123 |
| *AlternativeUserID* | *U* |  |
| *UserName* | *U* |  |
| *UserIsRequestor* | *U* |  |
| RoleIDCode | M | **EV( 121025, DCM, "Patient")** |
| Department |  | *"Admitting"* |
| Shift |  |  |
| NetworkAccessPointTypeCode | *U* |  |
| NetworkAccessPointID | *U* |  |

|  |  |  |  |
| --- | --- | --- | --- |
| **Participating Location Object**  (AuditMessage/ ParticipantObjectIdentification) | ParticipantObjectTypeCode | M | "2" (system object) |
| ParticipantObjectTypeCodeRole | M | **See section 6.x.5.4, location roles** |
| *ParticipantObjectDataLifeCycle* | *U* |  |
| *ParticipantObjectIDTypeCode* | *M* |  |
| *ParticipantObjectSensitivity* | *U* |  |
| ParticipantObjectID | *M* | *"Automatic admitting lobby 3"* |
| *ParticipantObjectName* | *U* |  |
| *ParticipantObjectQuery* | *U* |  |
| ParticipantObjectDetail | U |  |

{

Pri : "101",

version: "1",

Timestamp: "2015-03-17T00:15"

Hostname: "nemo.frodo.org"

App-name: "IHE+SOLE"

Procid: "1234"

Msg-id : "7c3fb1e7-e8e8-4df5-bde5-c224d3290a8b"

Msg : "lots of XML"

}

Appendices

*<Add any applicable appendices below; NA if none.>*

Appendix A – <Appendix A title>

Appendix A text goes here.

Appendix B – <Appendix B Title>

Appendix B text goes here.

B.1 <Add Title>

Appendix B.1 text goes here.

Volume 3 Namespace Additions

Add the following terms to the IHE Namespace:

<Please explicitly identify all new OIDs, UIDs, URNs, etc., defined specifically for this profile. These will be added to the IHE TF General Introduction namespace appendix when it becomes available. These items should be collected from the sections above by the author, and listed here as additions when this document is published for Trial Implementation. This section will be deleted prior to inclusion into the Technical Framework as Final Text, but should be present for publication of Public Comment and Trial Implementation.>

Volume 4 – National Extensions

Add appropriate Country section

4 National Extensions

None